## Sheet1
Problem | Possible reason | Solution
Problems with optical overlay | The area chosen for overlay has dirt particles | Choose an area with cleaner background
Change the slide
| High background noise and low contrast image | Change the imaging parameters to ensure proper contrast. Maintain a low exposure to ensure proper capturing of images while the cantilever is moving. In such case, we recommend normal brightfield imaging without polarizers in beam path for proper optical overlay. The polarizers can be put back in their place after opitcal overlay is completed.
| Cantilever is not in focus | Make sure the cantilever is in focus.
Incomplete or failed approach | Approach is too slow | Increase the approach speed, at the same time make sure that at the end of the approach the cantilever remains stable
Make sure that the bead is attached to center of the cantilever tip
| SetPoint value too low | Increase the SetPoint value by roughly 50%
| Laser alignment is not proper | Make sure laser is focused on the center of the cantilever tip
Spring constant (obtained after calibration) different than that declared by manufacturer | Too large or heavy bead | Switch to a smaller bead of light material
| Non specific interactions between bead and glass slide | Make sure the cantilever is in the solution freely (for the non-contact mode of calibration)
| Wrong peak selected in the thermal spectrum | Make sure the correct peak is selected (usually close to the resonance frequency value)
Force-Distance curves are irregular | Sample surface is extremely rough | Change sample. Follow the fixation timings precisely
| Deposition of salt due to evaporation of PBS | Make sure that PBS is not evaporated during the measurement.
| Stabilization unit not working properly | Restart the Stabilization unit
| Cantilever is not clean | Clean the cantilever extensively with 70% ethanol followed by distilled water. Change the cantilever if necessary.
| Repeated sampling of an area | Ensure that a single area is not measured more than once as the cantilever indentation has a viscoelastic effect on the sample surface
| Incorrect force or contact time | Optimize force and contact time for the sample by making sure that approach is successful and stable
| PBS/PFA is not clean | Prepare fresh solutions of PBS/PFA
| Calibration of cantilever is disturbed | Repeat calibration with input of concurrent temperature
| Bead is not properly atttached or detached during measurements | Use a new cantilever-bead ensemble
| Measurement on the edge of the sample | Use a different area of interest
Low laser voltage output | The surface of the cantilever holder is dirty | Clean the cantilever holder with non-scratching paper
| Misalignment of laser with the cantilever tip and detector | Realign the laser and repeat calibration
| Cantilever not mounted incorrectly mounted on the cantilever holder | Mount the cantilever centrally on the holder. Maintain a parallel distance throughout the length of cantilever and border of the holder
| Cantilever is defective | Change cantilever
| Adjust mirror position | Adjust the optical path by varying the mirror position in the AFM head
Highly deviant Young's modulus values | Incorrect model | Make sure tip-sample interaction is considered while selecting models
| Exclusion of data not performed properly | Exclude incorrect data
| Improper estimation of Contact point | Optimize analysis parameters for better determination of contact point