## Sheet1
Test | Single Generator (low activity) |  |  | Dual Generator (high activity)
| Validation #1 | Validation #2 | Validation #3 | Validation #1 | Validation #2 | Validation #3
Production yield [%] (non-decay corrected) | 70 | 65 | 65 | 69 | 74 | 68
Produced activity [GBq] | 0.63 | 0.61 | 0.57 | 1.54 | 1.71 | 1.55
Product volume [mL] (visually assessed) | 11 | 11 | 11 | 11 | 11 | 11
Appearance (clear, colourless) | Pass | Pass | Pass | Pass | Pass | Pass
pH [4-8] | 6.5 | 6.5 | 6.5 | 6.5 | 6.5 | 6.0
Half-life [62-74 min] | 71 | 66 | 74 | 69 | 68 | 67
Radionuclidic identity (MCA) [500-520 keV] | 516 | 517 | 516 | 514 | 512 | 510
Radionuclidic purity (MCA) [>99.9%] | Pass | Pass | Pass | Pass | Pass | Pass
Radiochemical purity (HPLC) [>95%] | 98.4 | 98.2 | 99.5 | 99.9 | 99.5 | 99.7
Radiochemical purity (iTLC) [>95%] | 98.1 | 97.0 | 96.6 | 99.7 | 99.3 | 99.6
Residual solvents (ethanol) [<79,000 mg/L] | 38,310 | 34,024 | 31,978 | 36,576 | 40,141 | 43,217
Bacterial endotoxin test [<17.5 EU/mL] | <10.0 | <10.0 | <10.0 | <10.0 | <10.0 | <10.0
Filter integrity test [>50 psi] | Pass | Pass | Pass | Pass | Pass | Pass
Radionuclidic purity (MCA) [<0.001% 68Ge-breakthrough] | Pass | Pass | Pass | Pass | Pass | Pass
Sterility test | Pass | Pass | Pass | Pass | Pass | Pass